SPECIFICATION CONTROL DRAWING

| Name: | Device Type: | Part Number: |
| --- | --- | --- |
| Thermopile Module | TPML1 | 2007 |

| Date |  | Sym | Rev. | Revision Record | Auth. | Dr. | Ch. |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 05JAN2000 |  | MS | 00 | Initial Release | MS | MS |  |
| 19JAN2000 |  | MS | 01 | Filter Transmission | MS | MS |  |
| 04FEB2000 |  | MS | 02 | Time Constant | MS | MS |  |
| 05APR2000 |  | MS | 03 | Test Tolerances | MS | MS |  |
| 28JUL2000 |  | MS | 04 | Drawing, Tolerances, Filter | MS | MS |  |
| Drawn | Mischa Schulze |  |  |  | Date | 28JUL2000 |  |
| Checked | Mischa Schulze |  |  |  | Date | 28JUL2000 |  |
| Approved |  |  |  |  | Date |  |  |
| Released |  |  |  |  | Date |  |  |

| Customer: |  | Reference: | First Used: |  |  |
| --- | --- | --- | --- | --- | --- |
| Checked |  |  |  | Date |  |
| Approved |  |  |  | Date |  |
| Released |  |  |  | Date |  |

1	SCOPE	3

1.1	PURPOSE	3

1.2	SCHEMATIC	3

2	GENERAL CHARACTERISTICS	3

2.1	ABSOLUTE MAXIMUM RATINGS	3

2.2	OPERATING CONDITIONS	4

2.3	HANDLING REQUIREMENTS	4

2.4	APPLICATION HINTS	4

2.4.1	Connection Requirements	4

2.4.2	Ambient Temperature Compensation	4

2.4.3	Measuring Tolerances	5

2.4.4	Applications Adjustment	6

3	TYPE CHARACTERISTICS	7

3.1	DESIGN CHARACTERISTICS	7

3.2	OPTICAL CHARACTERISTICS	7

3.2.1	Definition and Measurement	7

3.2.2	Horizontal and Vertical  Characteristics	8

3.3	FILTER CHARACTERISTICS	9

3.3.1	Parameter	9

3.3.2	Sample Curves	9

3.4	MECHANICAL REQUIREMENTS	9

3.4.1	Mechanical Dimensions of the Module	9

3.4.2	Mirror with Assembled Protection Filter	11

3.4.3	Scheme of Connection	11

3.4.4	Marking	11

3.5	ELECTRICAL REQUIREMENTS	12

3.5.1	Compensated Thermopile Output Voltage	12

3.5.2	Temperature Reference Output Voltage	13

4	TESTING	13

4.1	ELECTRICAL TEST	13

4.1.1	Test Conditions	13

4.1.2	Test Parameters	14

4.2	VISUAL TEST	14

4.2.1	Mirror Optics	14

4.2.2	Protection Filter	14

5	PACKAGING	14

5.1	PACKING FOR DELIVERY	14

5.2	LABELLING FOR DELIVERY	14

6	RELIABILITY	15

6.1	TEST CONDITIONS	15

6.2	ENVIRONMENTAL PERFORMANCE	15

6.3	MECHANICAL PERFORMANCE	15

7	QUALITY	16

7.1	QUALITY SYSTEM	16

7.2	LOT ACCEPTANCE TEST	16

7.2.1	Test Conditions and Parameter	16

7.2.2	Test Level	16

7.3	CHANGES AND MODIFICATIONS	16

7.4	REFERENCED DOCUMENTS	16

# SCOPE

## PURPOSE

This measurement system is used for the non-contact measurement of surface temperatures based on IR-radiation. The thermopile sensor is integrated in an ambient temperature compensated module with mirror or lens optics. An additional IR-window can be used for protection against contamination and condensation. The standard system is optimised to an object emission near 95%, but it can be modified to respective application conditions.

## SCHEMATIC

The following drawing shows the principle of the measurement and circuitry :

- Principle of the Measurement and Schematic Circuit Diagram

# GENERAL CHARACTERISTICS

## ABSOLUTE MAXIMUM RATINGS

| Supply Voltage (Single Power Supply) | +18V |
| --- | --- |
| Storage Temperature Range | -40°C to 105°C |

- Absolute Maximum Ratings

## OPERATING CONDITIONS

| Supply Voltage (Single Power Supply) | +5V .. +16V |
| --- | --- |
| Output Voltage Swing | 0V .. (Vsupply - 2V) |
| Supply Current (RL ) at V+ 5V | 1.5mA (typical) ; 5mA (maximum) |
| Power Supply Rejection Ratio | 50dB (minimum) |
| Output Current (Short Circuit to Ground) | 60mA (maximum) |
| Response Time | 20ms (typical) |
| Operating Ambient Range | -20°C .. 100°C |

- Operating Conditions

## HANDLING REQUIREMENTS

Stresses above the absolute maximum ratings may cause damages to the device.

Short circuits from the output to V+ can cause destruction's. Regarding to short circuits to ground, the maximum output current is approximately 60mA independent of V+. Continuous short-circuits at values of supply voltage in excess of +15VDC might cause destruction's.

Precautions should be taken to avoid reverse polarity of power supply. Reversed polarity of power supply results in a destroyed unit.

The module can be damaged by electrostatic discharges. Please take appropriate precautions for the handling.

Do not expose the sensor's to aggressive detergents such as freon, trichlorethylen, etc. Windows may be cleaned with alcohol and cotton swab.

Hand soldering may be applied by a maximum temperature of 260°C for a dwell time less than 10s.

## APPLICATION HINTS

### Connection Requirements

Capacitive loads which are applied directly to the outputs reduce the loop stability margin. Values of 50pF can be accommodated using the worst-case unity gain connection. A resistive isolation should be used if a larger load capacitance must be driven.

### Ambient Temperature Compensation

The thermopile sensor converts the temperature radiation of an object surface to an electrical signal by thermocouples (Seebeck effect). The sensor output voltage is caused by the temperature difference between radiation heated (hot) junctions and cold junctions with a good thermal contact to the housing. The changing of the housing (cold junction) temperature have to be detected to get the right output signal depending on the object radiation only. The sensor module is designed to compensate the ambient (sensor) temperature changing by a hardware adjustment. Because of many physical affects are influencing the non-contact temperature measurement based on infrared radiation, it is difficult to have the best initial adjustment for the different applications. Therefore some deviations could be found at first measuring. For all applications the optimised solution can be prepared and fixed based on the measurement in the application. Don't hesitate to contact the marketing or engineering department for support.

The temperature compensation is working right in an ambient temperature range, limited by different device parameters of the thermopile sensor and the thermistor. Please have a look to the following diagram for better understanding of the principle compensation curve. The central curve shows the deviation for an optimum working of a compensated module type 1. The other curves are showing the maximum deviation for the same module type at the specified object temperature (worst case).

The compensation of the mentioned module type is adjusted to the best fitting at 20°C to 60°C ambient temperature.

- Typical and Maximum Temperature Deviation of a Module Type 1

### Measuring Tolerances

The measuring tolerance of the sensor modules depends on the emissivity, object temperature, object size to spot size relation, temperature gradients on the sensor, device tolerances and the optimal adjustment of the ambient temperature compensation as basis for the measurement in a wide ambient range.

The following diagram shows the 'worst case' temperature tolerances at different object (black body) temperatures need to be added to the typical compensation curve for a standard module type 1.

- Maximum Temperature Deviation of a Module Type 1

### Applications Adjustment

Don’t hesitate to contact us, if the sensor modules show problems in your special application. The sensor module can be adjusted to the different requirements of most applications. If possible, please give following information's to allow us better support :

* Object temperature range 
* Object (surface) emissivity 
* Ambient temperature range
* Temperature accuracy of measurement at object and ambient temperature ranges
* Field of view (spot size and measuring distance)
* Sensor module with a special case
* Measuring data of the sensor module in your application (for the optimal adjustment of the ambient temperature compensation) like the following :
	
Tobject	Tambient	Vout(Tsensor)	      Vout(Tobject)
       :		       :		         :				    :		

The measuring conditions should be :
	- Constant object temperature
	- Constant object emissivity
	- Variation of the ambient (sensor) temperature in the used range
	- Measuring of the output voltages at different stabilised ambient temperatures

# TYPE CHARACTERISTICS

## DESIGN CHARACTERISTICS

| Parameter | Type | Drawing |
| --- | --- | --- |
| Thermopile Module | TPML |  |
| Thermopile Sensor | HEIMANN TPS434 | 2/70372 |
| Optics | HEIMANN Ellipsoid Mirror | 45501887 |
| Protection Filter | Type Standard Mounted |  |
| PCB | TPMF710 | 2/70709 |
| Schematic and Layout | TPMF710 | HW-WI-H80-S0007 |
| Spectral Range Protection Filter | HEIMANN STANDARD Coating |  |
| Spectral Range Sensor Filter | HEIMANN STANDARD Coating |  |

- Design Characteristics of the Specified Device Type

## OPTICAL CHARACTERISTICS

### Definition and Measurement

The field of view is defined by the full angle at the half energy points given by the following experiment.

- Measuring of the Field of View

### Horizontal and Vertical  Characteristics

- Horizontal Characteristics of the Mirror Optics

- Vertical Characteristics of the Mirror Optics

## FILTER CHARACTERISTICS

### Parameter

| Standard Filter |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Parameter | Limits |  |  | Units | Conditions |
|  | Min | Typ | Max |  |  |
| Average Transmission | 70 |  |  | % | Wavelength Range from 7.5µm to 13.5µm |
| Average Transmission |  |  | 0.5 | % | Wavelength Range from Visual to 5µm |
| Cut On | 5.2 | 5.5 | 5.8 | µm | At 25°C |

- Parameter of the Used Filter Type Standard

### Sample Curves

- Sample Filter Transmission Curves of the Standard Filter

## MECHANICAL REQUIREMENTS

### Mechanical Dimensions of the Module

- Dimensions of the Specified Module

### Mirror with Assembled Protection Filter

- Assembled Filter to Mirror of the Specified Module

### Scheme of Connection

- Scheme of Connection of the Specified Module

### Marking

Each device is marked at the backside of the mirror optics by a label with a letter and a serial number. The letter describes the place of manufacturing.

## ELECTRICAL REQUIREMENTS

### Compensated Thermopile Output Voltage

The compensated thermopile output voltage Vout(Tobject) is given by the circuit 4 of the pcb and circuit 1 of the connector.

The following table shows the typical and maximum tolerance voltages for different object temperatures (black body) at 25°C ambient (sensor) temperature :

| Black Body | Vout=f(Tobject) |  |  | Gradient |
| --- | --- | --- | --- | --- |
| Tobject | Minimum | Typical | Maximum | Typical |
| °C | V | V | V | V/ °C |
| 0 | 1.218 | 1.273 | 1.333 | 0.011 |
| 25 | 1.545 | 1.591 | 1.641 | 0.0145 |
| 50 | 1.945 | 2.004 | 2.066 | 0.0185 |
| 100 | 3.051 | 3.143 | 3.240 | 0.0272 |
| 200 | 6.581 | 6.783 | 6.990 | 0.0456 |

- Output Voltage Vout(Tobject) at different Object Temperatures

- The compensated thermopile output voltages are specified including the protection filter type standard and by the condition, that the whole target fills the field of view defined by the energy points below 1%.

| Parameter | Minimum | Typical | Maximum |
| --- | --- | --- | --- |
|  | ms | ms | ms |
| Time Constant | 5 | 20 | 100 |

- Response Time of the Thermopile Output

- The time constant can be measured as response to an object temperature 	   jump (low to high or high to low) based on the following equations :

Low to High : ΔV=ΔVmax∗1−e−tτt=τ		ΔV=ΔVmax∗1−1e

High to Low : ΔV=ΔVmax∗e−tτt=τ		ΔV=ΔVmax∗1e

### Temperature Reference Output Voltage

The temperature reference output voltage Vout(Tambient) is given by the circuit 1 of the pcb and circuit 4 of the connector.

The following table shows the typical and maximum tolerance voltages for different ambient temperatures :

|  | Vout=f(Tambient) |  |  |
| --- | --- | --- | --- |
| Tsensor | Minimum | Typical | Maximum |
| °C |  | V | V |
| -40 | 1.206 | 1.240 | 1.273 |
| -20 | 1.241 | 1.273 | 1.305 |
| 0 | 1.325 | 1.357 | 1.390 |
| 20 | 1.496 | 1.530 | 1.569 |
| 25 | 1.556 | 1.591 | 1.631 |
| 40 | 1.762 | 1.813 | 1.870 |
| 60 | 2.085 | 2.162 | 2.247 |
| 80 | 2.386 | 2.491 | 2.606 |
| 100 | 2.612 | 2.744 | 2.885 |

- Output Voltage Vout(Tambient) at different Ambient (Sensor) Temperatures

# TESTING

## ELECTRICAL TEST

### Test Conditions

| Object Size | > 50mm |
| --- | --- |
| Object Emission | > 99% |
| Object Temperature | 100°C ±0.5°C |
| Ambient Temperature | 25°C ±2°C |
| Viewing angle to the object | > 50° |

- Test Conditions for the Electrical Parameters

### Test Parameters

| Typical | Typical | Vout=f(Tobject) |  |  | Vout=f(Tambient) |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Tobject | Tambient | Minimum | Typical | Maximum | Minimum | Typical | Maximum |
| °C | °C | V | V | V | V | V | V |
| 100 | 23.0 | 3.053 | 3.147 | 3.243 | 1.531 | 1.566 | 1.605 |
| 100 | 23.5 | 3.052 | 3.146 | 3.242 | 1.537 | 1.572 | 1.611 |
| 100 | 24.0 | 3.053 | 3.145 | 3.242 | 1.543 | 1.578 | 1.618 |
| 100 | 24.5 | 3.052 | 3.144 | 3.241 | 1.549 | 1.585 | 1.624 |
| 100 | 25.0 | 3.051 | 3.143 | 3.240 | 1.556 | 1.591 | 1.631 |
| 100 | 25.5 | 3.050 | 3.142 | 3.239 | 1.562 | 1.598 | 1.637 |
| 100 | 26.0 | 3.049 | 3.142 | 3.238 | 1.568 | 1.604 | 1.644 |
| 100 | 26.5 | 3.047 | 3.141 | 3.238 | 1.574 | 1.611 | 1.652 |
| 100 | 27.0 | 3.047 | 3.140 | 3.238 | 1.580 | 1.617 | 1.659 |

- Test Parameters for the Electrical Device Test

- The compensated thermopile output voltages are specified including the mounted protection filter type standard and by the condition, that the whole target fills the field of view defined by the energy points below 1%.

## VISUAL TEST

### Mirror Optics

The visual inspection of the mirror optics has to be performed according to the document number S0001 index 1, component specification of the sensor housing.

### Protection Filter

The visual inspection of the protection filter has to be performed by a visible light reflection test.

# PACKAGING

## PACKING FOR DELIVERY

The packing for the delivery of mass production devices depends on the respective order quantity.

## LABELLING FOR DELIVERY

Each label contains the following information :

Name of the manufacturer, Product Group, Product type, Product number,

Quantity per box, Date of packing, Place of packing

# RELIABILITY

HEIMANN confirms that samples of modules have been tested to the following reliability conditions.

## TEST CONDITIONS

| Object Size | > 50mm |
| --- | --- |
| Object Emission | > 99% |
| Object Temperature | 95°C ±0.3°C |
| Ambient Temperature | 25°C ±2°C |
| Viewing angle to the object | > 50° |

- Test Conditions for the Electrical Parameters

## ENVIRONMENTAL PERFORMANCE

| Parameter | Test Condition | Test Result |
| --- | --- | --- |
| Heat Ageing | (1005)°C, 1000hours, Operation V+16V | Vout(Tobject) Changing  2% |
| Cold Ageing | (-205)°C, 1000hours, Operation V+16V | Vout(Tobject) Changing  2% |
| Temperature Cycling | (-20±5)°C to (85±5)°C, 2hours/each, 30cycles, Operation V+16V | Vout(Tobject) Changing  2% |
| Humidity | (60±5)°C, (80±5)% rH, 1000hours, Operation V+16V | Vout(Tobject) Changing  2% |

- Reliability Test Results for the Environmental Performance

## MECHANICAL PERFORMANCE

| Parameter | Test Condition |  | Test Result |
| --- | --- | --- | --- |
| Mechanical Shock | 1m free fall onto wooden plate, x-y-z-directions |  | Vout(Tobject) Changing  2% |
| Vibration | Acceleration 20g 50g 100g 500g 1000g 5000g 10000g | Pulse Duration 3ms 3ms 3ms 1ms 0.7ms 0.15ms 0.08ms | Vout(Tobject) Changing  2% Vout(Tobject) Changing  2% Vout(Tobject) Changing  2% Vout(Tobject) Changing  2% Vout(Tobject) Changing  2% Vout(Tobject) Changing  2% Vout(Tobject) Changing  2% |

- Reliability Test Results for the Mechanical Performance

# QUALITY

## QUALITY SYSTEM

HEIMANN is an ISO 9001 certified manufacturer with established SPC and TQM. All materials are checked according to specifications and final goods meet the specified tests.

## LOT ACCEPTANCE TEST

### Test Conditions and Parameter

According to item 4.1.1 and 4.1.2 .

### Test Level

Lot conformance to specification of products delivered in volume production is checked by means of following tests :

| Test | Conditions | Level |
| --- | --- | --- |
| Electrical parametric Vout=f(Tobject) | Acc. to the test parameters at the test conditions , tolerance check | 100% |
| Electrical parametric Vout=f(Tambient) | Acc. to the test parameters at room temperature , functional check | 100% |
| External visual | Physical correctness | 0.65 |
| Mechanical | External dimensions | 0.65 |

- Acceptable Quality Level for the Lot Acceptance Tests

## CHANGES AND MODIFICATIONS

Changes or modifications at the product which could influence performance and/or quality of the device have to be announced to said customer in advance and approved by said customer.

## REFERENCED DOCUMENTS

The referenced documents form a part of this drawing.  The revision level of these referenced documents unless defined shall be that  which is in effect on the date of the purchase order.